## Slide 1: Op-Amp ECG Heartbeat Detector
Precision rectifier · peak detector · Schmitt comparator
Brahiam Alberto Campiño Romero
Instructor: Eng. Rodolfo José Gutiérrez González
Nonlinear Circuits · Universidad de Ibagué · 2026
Notes: Heartbeat detector overview: the goal is to deliver one clean pulse per QRS complex using only operational amplifiers.

## Slide 2: Problem
Design and implement an op-amp-based heartbeat detector powered from ±12 V, able to identify every QRS complex of an ideal 0.5 to 2 Vpp ECG-like signal regardless of its polarity or waveform type, delivering one clean rectangular pulse per beat over a range of 40 to 150 beats per minute.

## Slide 3: One pulse per beat, regardless of polarity
Amplifier
×4.9
Rectifier
|Vg|
Peak detector
τ = 100 ms
Schmitt
1.01 / 0.52 V
LED
output
Nonlinear blocks
Why a linear circuit is not enough
The R wave can be positive or negative → rectify
The QRS has several lobes → merge them into one
Design limits: 0.5 to 2 Vpp
Notes: Each block solves one problem: the rectifier removes the polarity dependence, the peak detector merges the QRS lobes into a single event, and the Schmitt trigger turns the envelope into a clean digital pulse.

## Slide 4: Full circuit: 6 op-amps at ±12 V
Simulated in Proteus with LM741
Notes: Top row: amplifier, half-wave rectifier and summer. Bottom row: buffered peak detector, Schmitt comparator and LED. The oscilloscope probes input, Vr, Venv and output.

## Slide 5: Stage 1 · Non-inverting amplifier
1
Key points
Vg = Vin (1 + Rf1/Rg1)
Av = 4.9
Rg1 = 10 kΩ → Rf1 = 39 kΩ
0.5–2 Vpp → 2.45–9.8 V
Av = 4.9
valid range: 4.04 to 5.25
Notes: The gain must lie between 4.04 (to clear the threshold with margin) and 5.25 (to avoid saturation). 39 k over 10 k gives 4.9 with standard values.

## Slide 7: Stage 2 · Precision rectifier
2
Vg > 0
D2 conducts · Vhw = −Vg
Vg < 0
D1 conducts · Vhw = 0
Precision
Dead zone ≈ 3 µV
Notes: The diode is inside the feedback loop: the op-amp drives its output an extra 0.6 V, so that drop does not appear at Vhw. D1 keeps the op-amp out of saturation on the negative half-cycle.

## Slide 8
Case 1, Vg +

## Slide 9
Case 2, Vg -

## Slide 10: Stage 2 · Summer: from half-wave to |Vg|
2
Key points
Vr = −(Vg + 2·Vhw)
R6 = R4 = 20 kΩ
R6 = 2·R5 (R5 = 10 kΩ)
Result: Vr = |Vg|
1 %
tolerance for R2 to R6
Notes: R4 feeds the full signal with weight 1 and R5 the half-wave with weight 2. On the positive half-cycle this flips the sign; on the negative one only R4 acts. Resistors must be 1 %.

## Slide 14: Stage 3 · Peak detector
3
Key points
τ = R9·C1 = 100 kΩ · 1 µF = 100 ms
Valid range: 90–110 ms
Buffer: τ is independent of R8
100 ms
discharge time constant
0.1 ms
charge time constant
Notes: τ above 90 ms so the beat is not released between QRS lobes; below 110 ms so it releases before the next beat at 150 bpm. The buffer keeps R8 from discharging the capacitor.

## Slide 19: τ sets the maximum heart rate
Chart: Input amplitude (Vpp)
| 0.3 | 0.4 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 1.0 | 1.1 | 1.2 | 1.3 | 1.4 | 1.5 | 1.6 | 1.7 | 1.8 | 1.9 | 2.0
τ = 68 ms | 413 | 364 | 333 | 312 | 296 | 283 | 273 | 264 | 257 | 250 | 245 | 240 | 236 | 232 | 228 | 225 | 222 | 219
τ = 100 ms (design) | 336 | 290 | 262 | 242 | 228 | 217 | 208 | 201 | 195 | 189 | 185 | 181 | 177 | 174 | 171 | 168 | 165 | 163
τ = 150 ms | 261 | 220 | 196 | 180 | 168 | 159 | 152 | 146 | 141 | 137 | 133 | 130 | 127 | 125 | 122 | 120 | 118 | 117
150 bpm target | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150
Formula (taken from the report)
163 bpm
limit at 2 Vpp
Notes: A larger peak takes longer to discharge, so the limit drops with amplitude. With τ = 100 ms the worst case (2 Vpp) is 163 bpm, above the 150 bpm target. Simulation confirmed that at 180 bpm with 2 Vpp the comparator stays high.

## Slide 21: Stage 4 · Schmitt comparator
4
Key points
k = R10/R11 = 0.0233 (430 kΩ)
Vref = 0.75 V (R12 = 15 kΩ, R13 = 1 kΩ)
1.01 V
VTH (trigger)
0.52 V
VTL (release)
0.49 V
hysteresis
Notes: Positive feedback: there is no virtual short. VTH sits below the minimum peak (1.22 V) and VTL well above the offset. Hysteresis prevents noise-induced chatter.

## Slide 25: Proteus simulation
1 V pulse · 0.5 Hz
2 Vpp sine · 0.5 Hz
Synthetic heartbeat
Yellow: input · Blue: Vr · Pink: Venv · Green: output
Notes: Pulse: the envelope holds and decays; the output lasts about 400 ms. Sine: two pulses per cycle, showing both polarities are detected. Synthetic heartbeat: a single pulse per beat despite multiple lobes and the T wave.

## Slide 27: Simulation vs calculation
4.90 V
Vr peak (calc. 4.90 V)
400 ms
output width (calc. 405 ms)
2
pulses per sine cycle
≤ 1.6 %
maximum error
Pending in the lab
Measure τ, VTH, VTL and offset on the real build
Notes: Proteus values match the calculations with errors of 1.6 % or less. Comparison with bench measurements is still pending.

## Slide 28: Conclusions
1
Polarity The precision rectifier detects both positive and negative R waves
2
One pulse The peak detector merges the QRS; τ = 100 ms is the critical parameter
3
No chatter Schmitt hysteresis gives clean transitions
4
Validated Proteus matches the calculation (error ≤ 1.6 %)
Notes: Possible improvements: instrumentation amplifier with band-pass filter, adaptive threshold, and beat counting with a microcontroller.

## Slide 29: References https://github.com/Brahiam18/ecg-heartbeat-detector/tree/main